GMP G3 FESA Interface
Property name	Property / type	Field name	Field type	unit	description
Status	status
		status-item	GSI-status-field		OK/ERROR
		detailed-status-field	GSI-Detailed-status-field		0 – Power supply fault / 1 – Microwave head fault / 2 – Miscellaneous fault
		mircrowaveHeadStatus	bitset	/	0 – Head Fan / 1 – Microwave head interlock 1 / 2 – Filament fuse / 3 – Filament Fault / 4 – Arc detection / 5 – Magnetron water flow / 6 – Magnetron temperature / 7 – High voltage connector / (0=OK ; 1=FAULT)
		powerSupplyStatus	bitset	/	0 – DriverUnderVoltage / 1 – Heatsink overheat / 2 – Power supply Fan / 3 – Under output power / 4 – Power supply under voltage / 5 – Power supply over voltage / 6 - Power supply primary over current / 7 – Power supply primary over current / 8 - High voltage connector / 9 - Low voltage 24V / 10 - RP disjunction / 11 - External communication / 12 - Internal communication / 13 - Stop from front panel / 14 - Power supply water flow / 15 - Over output power / (0=OK ; 1=FAULT)
		systemStatus	bitset	/	0 - External security / 1 - RP limitation detection / 2 - Menu is activated on front panel / 3 - 0 / 4 - 0 / 5 - One fault on power supply / 6 - One fault on power supply / 7 - One fault on head / 8-15 - / (0=OK ; 1=FAULT)
Power	power				Turns device ON/OFF
Reset	reset				Resets device.
Init	init				?
Version	version				Returns generator reference. Fesa version. Class version. ?

Setting	setting	power	int	W	Power setpoint (FP setpoint)

Acquisition	acquisition	power	int	W	Actual power (FP readback)
		reflectionPower	int	W	Actual reflection power (RP readback)
		setPower	int	W	Users power setpoint
		generatorState	bool	/	0=Not ready / 1=Ready
		microwaveState	bool	/	0=OFF / 1=ON
		filamentState	bool	/	0=Fault or Off or Pre heating / 1=Ready
		mircrowaveHeadStatus	bitset	/	0 – Head Fan / 1 – Microwave head interlock 1 / 2 – Filament fuse / 3 – Filament Fault / 4 – Arc detection / 5 – Magnetron water flow / 6 – Magnetron temperature / 7 – High voltage connector / (0=OK ; 1=FAULT)
		powerSupplyStatus	bitset	/	0 – DriverUnderVoltage / 1 – Heatsink overheat / 2 – Power supply Fan / 3 – Under output power / 4 – Power supply under voltage / 5 – Power supply over voltage / 6 - Power supply primary over current / 7 – Power supply primary over current / 8 - High voltage connector / 9 - Low voltage 24V / 10 - RP disjunction / 11 - External communication / 12 - Internal communication / 13 - Stop from front panel / 14 - Power supply water flow / 15 - Over output power / (0=OK ; 1=FAULT)
		systemStatus	bitset	/	0 - External security / 1 - RP limitation detection / 2 - Menu is activated on front panel / 3 - / 4 - / 5 - One fault on power supply / 6 - One fault on power supply / 7 - One fault on head / 8-15 - / (0=OK ; 1=FAULT)
MicrowaveSwitch	setting	action	bool	/	True = Turn microwave ON / False = Turn microwave OFF
FilamentSwitch	setting	action	bool	/	True = Turn filament ON / False = Turn filament OFF
PulseOnOptical	command	/	/	/	Pulse on optical fibre
PulseInternal	command	/	/	/	Internal pulse

Configuration	setting	frequency	int	Hz	Frequency setpoint
		pulseLength	float	ms	Pulse length setpoint
		maxReflectionPower	int	W	Maximum RP setpoint

ExpertConfiguration	setting	startModes	enum	/	Select one of Normal, Overshoot, Ramp
		dutyCycle	int	%	Duty cycle setpoint
		highPulseLevel	int	W	Pulse high level setpoint
		lowPulseLevel	int	W	Pulse low level setpoint
		cwMode	bool	/	True=On starting the generator is in CW mode / False=On starting the generator is not in CW mode
		rampDuration	int	s	Adjust ramp duration